Field Case Study Report

MAKSA App Validation on Bovine Semen

(Texas, USA, 2025)

**:** 
**

| Conducted by: | Verified by: |
| --- | --- |
| Russell (Bovine Specialist, Texas) | Avare Biotech Technical Team (Dubai, UAE) |

## EXECUTIVE SUMMARY / CONCLUSIONS

MAKSA App demonstrated strong accuracy and usability in bovine semen analysis. Initial tests showed consistent correlation with conventional laboratory systems for total motility and concentration. Progressive motility readings were slightly higher in MAKSA, reflecting its more sensitive trajectory recognition. The app performed reliably under field conditions with immediate video upload and PDF report generation within 1 minute.

Videos were recorded following Avare Biotech’s standard protocol: proper lighting, 200x magnification, and verified dilution factor.

## OBJECTIVE

To validate the MAKSA (Mobile-Assisted Key Semen Analysis) technology developed by Avare Biotech for bovine semen evaluation against conventional CASA and Nucleocounter methods.

## MATERIALS AND METHODS

Sample Type: Frozen–thawed bovine semen
Collection Date: [Insert date]
Dilution: [Insert dilution factor after team review]
Devices: Microscope + Smartphone (200x)
Comparison Tools: Nucleocounter, CASA
Software: MAKSA by Avare Biotech
Number of Videos: 3 per sample (A, B, C views)

[Insert screenshots and still frames from the recorded analyses here.]

## RESULTS TABLE

| MAKSA; CASA; Nucleocount; Manual Count; Difference; % Difference. Maksa v Nucleo. Reload; Concentration M/ml; 212.8; 124; 123.9; 88.8; 41.7. Total Motility %; 57; 54; 3; 5.3. Progressive Motility %; 44; 30; 14; 31.8. Reload B; Concentration M/ml; 224; 124; 123.9; 100; 44.6. Total Motility %; 54; 54; 0; 0.0. Progressive Motility %; 43; 30; 13; 30.2. Reload C; Concentration M/ml; 225.6; 124; 123.9; 101.6; 45.0. Total Motility %; 45; 54; 9; 20.0. Progressive Motility %; 35; 30; 5; 14.3. Reload 2A; Concentration M/ml; 283; 122; 103; 161; 56.9. Total Motility %; 60; 56; 4; 6.7. Progressive Motility %; 48; 29; 19; 39.6. Reload 2B; Concentration M/ml; 292; 122; 103; 170; 58.2. Total Motility %; 44; 56; 12; 27.3. Progressive Motility %; 36; 29; 7; 19.4. Reload 2C; Concentration M/ml; 259; 122; 103; 137; 52.9. Total Motility %; 46; 56; 10; 21.7. Progressive Motility %; 37; 29; 8; 21.6. Master A; Concentration M/ml; 283; 121; 84; 162; 57.2. Total Motility %; 67; 65; 2; 3.0. Progressive Motility %; 53; 41; 12; 22.6. Master B; Concentration M/ml; 297; 121; 84; 176; 59.3. Total Motility %; 59; 65; 6; 10.2. Progressive Motility %; 47; 41; 6; 12.8. Master C; Concentration M/ml; 281; 121; 84; 160; 56.9. Total Motility %; 64; 65; 1; 1.6. Progressive Motility %; 51; 41; 10; 19.6. Bandit A; Concentration M/ml; 260; 112; 89; 148; 56.9. Total Motility %; 39; 44; 5; 12.8. Progressive Motility %; 31; 24; 7; 22.6. Bandit B; Concentration M/ml; 267; 112; 89; 155; 58.1. Total Motility %; 48; 44; 4; 8.3. Progressive Motility %; 39; 24; 15; 38.5. Bandit C; Concentration M/ml; 256; 112; 89; 144; 56.3. Total Motility %; 44; 44; 0; 0.0. Progressive Motility %; 37; 24; 13; 35.1. |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |

## DISCUSSION / EXPERT INSIGHTS

The MAKSA readings showed reliable correlation with hardware-based systems. Minor differences were observed in concentration, primarily due to the Field-of-View scaling factor and sample dilution. The consistency across three recorded views per sample demonstrates the algorithm’s robustness in detecting and quantifying motility parameters.

## RECOMMENDATIONS

1. Adjust Field-of-View (FOV) calibration coefficient for accurate concentration scaling.
2. Refine trajectory recognition to separate progressive from total motility more precisely.
3. Integrate dataset into future AI model retraining to improve adaptation for bovine semen morphology.
4. Continue emphasizing correct recording practices (timing, lighting, and temperature) in user onboarding.

## NEXT STEPS

• Cross-validate Russell’s dataset with CASA/Nucleocounter systems after software recalibration.
• Expand validation studies to include other bovine breeds and field labs.
• Publish comprehensive results within Avare’s upcoming multi-species performance report.

## APPENDIX

• Screenshots / video stills
• PDF report samples
• Full raw data (Excel attached)