# Evaluating MAKSA Mobile App's Accuracy in Assessing Motility of Fresh Bovine Semen: A Comparative Study with Computer-Assisted Sperm Analysis (CASA)

# (Okara, Pakistan 2026)

Date: February 2026

Location: Okara, Pakistan

| Conducted by: Dr. Misbah Aslam (Okara, Pakistan) | Verified by: Avare Biotech Team (Dubai, UAE) |
| --- | --- |
| Lab technologist at Sexed sorted Semen Lab (Pakistan) I'm a vet specialized in animal reproduction. | Maksim Mershiev, The CEO&Co-Founder of Avare Biotech |
| Signature: | Signature: |

## Executive Summary / Conclusions

## Smart phone based MAKSA technology demonstrated an ease in Fesh Bovine semen Evaluation.

## Objective

To validate the smart phone based MAKSA (Mobile-Assisted Key Semen Analysis) technology  for Fresh Bovine Semen Analysis, showing good agreement with CASA (Androvision miniTube) for total motility, progressive motility, and concentration (NucleoCounter), indicating MAKSA's potential as a reliable on-farm tool.

## Materials and Methods

| Table 1. MAKSA results comparison |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Tests done | Key Semen Parameters | MAKSA | CASA Motility | NucleoCounter Conc. | Difference % |
| Bull 1 (Auburn Rojo) | Concentration (M/ml) | 1740.7 |  | 1356 | -22 |
|  | Total Motility (%) | 84 | 71.22 |  | -15.21 |
|  | Progressive Motility (%) | 72 | 70.39 |  | -2.23 |
| Bull 2 (Braford) | Concentration (M/ml) | 921.2 |  | 1002 | 8.77 |
|  | Total Motility (%) | 79 | 82.16 |  | 4 |
|  | Progressive Motility (%) | 75 | 82 |  | 9.33 |
| Bull 3 (Gabru) | Concentration (M/ml) | 464.1 |  | 389.5 | -16.07 |
|  | Total Motility (%) | 82 | 68.67 |  | -16.25 |
|  | Progressive Motility (%) | 73 | 68.63 |  | -5.98 |
| Bull 4 (SK-1 Gulabi ) | Concentration (M/ml) | 1162.2 |  | 912 | -21.52 |
|  | Total Motility (%) | 80 | 77.38 |  | -3.27 |
|  | Progressive Motility (%) | 76 | 76.4 |  | 0.52 |
| Bull 5 (Janos Manso) | Concentration (M/ml) | 1420.7 |  | 921 | -35.17 |
|  | Total Motility (%) | 81 | 69.74 |  | -13.90 |
|  | Progressive Motility (%) | 71 | 69.74 |  | -1.76 |
| Bull 6 (Lotto 062M) | Concentration (M/ml) | 464.1 |  | 405 | -12.73 |
|  | Total Motility (%) | 82 | 72.03 |  | -12.15 |
|  | Progressive Motility (%) | 73 | 71.91 |  | -1.49 |

## Discussion / Expert Insights

- Total and progressive motility showed strong correlation between MAKSA and CASA readings.

- Concentration values differed between MAKSA and CASA, suggesting need for app calibration with fresh-diluted semen data.

- MAKSA is a valuable on-farm tool for quick semen analysis, with potential for improved accuracy with calibration enhancements.

## Recommendations

- Add in-app calibration videos for fresh semen analysis.

- Include training mode with feedback for new users.

- Share field condition tips in the app for on-farm use.

## Appendix

• Screenshot of the semen videos made and CASA results
• MAKSA PDF report samples

| ## Bull 1: CASA readings | ## Bull 1: MAKSA values |
| --- | --- |
| ## Bull 2: CASA readings | ## Bull 2: MAKSA values |
| ## Bull 3: CASA readings | ## Bull 3: MAKSA values |
| ## Bull 4: CASA readings | ## Bull 4: MAKSA values |
| ## Bull 5 : CASA readings | ## Bull 5: MAKSA values |
| ## Bull 6: CASA readings | ## Bull 6: MAKSA values |