Frozen-thawed/Fresh

semen quality analysis (QA)

by Avare Biotech App protocol

Material: frozen–thawed bovine/equine/ovine semen (3-5 straws from different animals)

Goal: to verify that the app correctly measures semen concentration, total motility, and progressive motility compared to CASA and manual counting.

Tasks

| 1. Assess three key parameters: | 2. Compare results between: |
| --- | --- |
| - Concentration - Total motility - Progressive motility | - Avare App (MAKSA technology) - CASA (e.g., AndroScope, Isperm, etc.) - Manual microscope counting |

Steps

Thawing:

- Place a straw in 37 °C water bath for 30–45 sec, dry, gently shake.

- Work quickly: insemination should be performed within ≤10 min after thaw.

Preparation:
- Pre-warm a normal lab or a CASA slide (if you are rich enough) on a heating stage (37–38 °C).

Sampling:

- Check your cover slip size

- Choose the semen volume based on the table below

| Standard Cover-slip Dimensions, mm | Semen Volume, μl |
| --- | --- |
| 10x10 | 2 |
| 12x12 | 3 |
| 15x15 | 4.5 |
| 18x18 | 6.5 |
| 20x20 | 8 |
| 22x22 | 10 |

- Pipette only recommended volume µL from the straw and load on the slide (avoid bubbles) using a Microliter Volume Pipettor

Testing:

- Record 5 videos while moving the slide via microscope according to the instruction of Avare Biotech

- Here is the link to 1 minute video: https://www.youtube.com/shorts/CdFFHB0CZUg

- Run the same sample through CASA system on CASA slides (for accuracy if available on site).
Count sperm manually under microscope.

Data Recording:
- Enter results for all three methods to a spreadsheet.
- Note agreements or discrepancies.

Conclusion:
- ±10–15% difference = acceptable validation.
- >15% deviation — attach photo/video evidence and repeat with a new slide.

## Example Report

| Sample (Bull/Straw) | Parameters/Method | Avare App powered by MAKSA tech | CASA | Manual Count | Difference, % | Notes |
| --- | --- | --- | --- | --- | --- | --- |
| Bull A – Straw 1 | Concentration (M/mL) |  |  |  |  |  |
|  | Total Motility (%) |  |  |  |  |  |
|  | Progressive Motility (%) |  |  |  |  |  |

Practical Note:

- Using 10 µL from a 0.25–0.5 mL straw does not prevent insemination if done immediately.

- Best practice in AI stations is to use the same straw taking one drop for Avare analysis and inseminating the rest (90%) right away.